Enzymatic Assay of THYMIDINE PHOSPHORYLASE

(EC 2.4.2.4)

PRINCIPLE:

Thymidine + Orthophosphate  Thymine + 2-Deoxy-R-P

Abbreviations used:

TP = Thymidine Phosphorylase

2-Deoxy-R-P = 2-Deoxy-D-Ribose 1-Phosphate

CONDITIONS: T = 25C, pH = 7,4, A290nm, Light path = 1 cm

METHOD: Continuous Spectrophotometric Rate Determination

REAGENTS:

- 200 mM Potassium Phosphate Buffer, pH 7,4 at 25C.
(Prepare 500 ml in deionized water using Potassium Phosphate, Monobasic, Anhydrous. Adjust to pH 7,4 at 25C with 1 M NaOH.)

- 1 mM Thymidine Solution
(Prepare 100 ml in Reagent A using Thymidine.)

- 10 mM Potassium Phosphate Buffer, pH 7,0 at 25C (Enzyme Diluent)
(Prepare 100 ml in deionized water using Potassium Phosphate, Monobasic, Anhydrous. Adjust to pH 7,0 at 25C with 1 M NaOH.)

- Thymidine Phosphorylase Enzyme Solution
(Immediately before use, prepare a solution containing 1 - 2 units/ml of Thymidine Phosphorylase in cold Reagent C.)

PROCEDURE:

Pipette (in milliliters) the following reagents into suitable quartz cuvettes:

|  | Test | Blank |
| --- | --- | --- |
| Reagent B (Thymidine) | 3.00 | 3.00 |

Equilibrate to 25C. Monitor the A290nm until constant1, using a suitably thermostatted spectrophotometer. Then add:

| Reagent C (Enzyme Diluent) | ------ | 0.03 |
| --- | --- | --- |
| Reagent D (Enzyme Solution) | 0.03 | ------ |

Immediately mix by inversion and record the decrease in A290nm for approximately 5 minutes. Obtain

the A 290nm/minute using the maximum linear rate for both the Test and Blank.

CALCULATIONS:

Units/ml enzyme = (ΔA290nm/min Test - ΔA290nm/min Blank)(3,03)(df)(1,0)(0,03)

3,03 = Total volume (in milliliters) of assay

df = Dilution factor

1,0 = Difference in the millimolar extinction coefficient between thymidine and thymine under the

assay conditions

0,03 = Volume (in milliliters) of enzyme used

UNIT DEFINITION2:

One unit will convert 1.0 μmole each of thymidine and phosphate to thymine and 2-deoxyribose 1-phosphate per minute at pH 7.4 at 25C.

FINAL ASSAY CONCENTRATIONS:

In a 3,03 ml reaction mix, the final concentrations are 198 mM potassium phosphate, 1 mM thymidine and 0,03 – 0,06 unit thymidine phosphorylase.

REFERENCE:

Krenitsky, T.A. and Bushby, S.R.M. (1979) United States Patent 4,178,212, 1-8, Burroughs

Wellcome Co., Research Triangle Park, NC

NOTES:

1. The initial A290nm should be between 1,8 and 1,95. If the absorbance is not within the specific range, then fresh reagents should be prepared.

2. One unit is equivalent to 0,614 International Unit.

3. This assay is based on the cited reference.